## dilutions MFC_ChS
| MFC 1 DDE - pure cylinder - N2 matrix |  |  | MFC 2 DDE2 : cylinder zero air |  |  |  | MFC 1 DDE - pure cylinder - N2 matrix
Range mL/min | 0-10 |  |  | 0-1000 |  |  |  | 0-200mL/min
| % setpoint | Flow MFM measured (T=20°C, P=1013.25hPa) |  | % setpoint | Flow MFM measured (T=20°C, P=1013.25hPa) |  |  | % setpoint | Flow MFM measured (T=20°C, P=1013.25hPa)
| 10 | 1.04 |  | 10 | 101 |  |  | 10 | 20.4
| 20 | 2.04 |  | 20 | 203 |  |  | 20 | 40.5
| 30 | 3.04 |  | 30 | 304 |  |  | 30 | 60.5
| 40 | 4.04 |  | 40 | 405 |  |  | 40 | 80.7
| 50 | 5.04 |  | 50 | 505 |  |  | 3 | 6.1
| 60 | 6.04 |  | 60 | 606 |  |  | 5 | 10.3
| 70 | 7.04 |  | 70 | 706 |  |  | 7.4 | 15.1
| 80 | 8.05 |  | 80 | 806 |  |  | 15.8 | 32
| 90 | 9.05 |  | 90 | 906 |  |  | 20.8 | 42
| 100 | 10.05 |  | 100 | 1007 |  |  | 29.8 | 60
|  |  |  | 3 | 30
|  |  |  | 5 | 51
|  |  |  | 8 | 81
|  |  |  | 6.8 | 69
|  |  |  | 49.1 | 494
|  |  |  | 5.8 | 58
|  |  |  | 4.1 | 42
|  |  |  | 2.9 | 30
|  |  |  | 1.9 | 20
| y=a.x+b |  |  | y=a.x+b |  |  |  | y=a.x+b
| a= | 0.10013 |  | a= | 10.0694 |  |  | a= | 2.011
| b= | 0.036 |  | b= | 0.72838 |  |  | b= | 0.20047
Linearisation not used because flow fixed during all the test
Concentrations awaited after dilution :
|  |  | Point 1 | Point 2 | Point 3 | Point 4 | Point 5 | Point 6 | TEST
| MFC 0-200mL/min | Set Point MFC N2 | 8 | 8 | 16 | 20 | 35 | 34
| MFC 0-1000mL/min | Set Point MFC Air to dilute | 49 | 6.8 | 5.7 | 4 | 3.1 | 4
|  | N2 flow mL/min - pure cyl | 16.2883 | 16.2883 | 32.3761 | 40.42 | 70.5846 | 68.5736 | 70
|  | Air flow for dilution mL/min | 494.1275 | 69.2001 | 58.1238 | 41.0059 | 31.9434 | 41.0059 | 50
|  | Total flow mL/min | 510.4158 | 85.4884 | 90.4999 | 81.4258 | 102.528 | 109.5795 | 120
|  | C pure cylinder (ppbv) | C out (ppb) | C out (ppb) | C out (ppb) | C out (ppb) | C out (ppb) | C out (ppb) | C out (ppb)
Bouteille n°1 | DES | 476 | 15.19 | 90.6932 | 170.2877 | 236.2875 | 327.6984 | 297.8755 | 277.6667
| SO2 | 484 | 15.4453 | 92.2175 | 173.1497 | 240.2587 | 333.2059 | 302.8818 | 282.3333
| MeSH | 524 | 16.7218 | 99.8388 | 187.4595 | 260.1148 | 360.7436 | 327.9134 | 305.6667
| EtSH | 478 | 15.2538 | 91.0743 | 171.0032 | 237.2803 | 329.0753 | 299.1271 | 278.8333
| DMS | 484 | 15.4453 | 92.2175 | 173.1497 | 240.2587 | 333.2059 | 302.8818 | 282.3333
| CS2 | 488 | 15.5729 | 92.9796 | 174.5806 | 242.2444 | 335.9597 | 305.385 | 284.6667
Bouteille n°2 | DMDS | 502 | 16.0197 | 95.6471 | 179.5891 | 249.194 | 345.5979 | 314.146 | 292.8333
Bouteille n°3 | H2S | 500 | 15.9559 | 95.266 | 178.8736 | 248.2012 | 344.221 | 312.8945 | 291.6667
| cycles repetitions (number) |  | 8 | 12 | 8 | 8 | 10
| duration of one cycle (s) |  | 600 | 600 | 600 | 600 | 600
| duration for one dilution point (s) |  | 4800 | 7200 | 4800 | 4800 | 6000
| total duration (s) | 27600
| total duration (hours) | 7:40 | 1:20 | 2:00 | 1:20 | 1:20 | 1:40
Bouteille COS INERIS
Concentrations awaited after dilution :
|  |  | Point 1 | Point 2 | Point 3 | Point 4 | Point 5 | Point 6 | TEST
| MFC 0-10mL/min | Set Point MFC N2 | 3 | 4 | 4.7 | 5.5 | 7
| MFC 0-1000mL/min | Set Point MFC Air to dilute | 100 | 100 | 100 | 100 | 100
|  | N2 flow mL/min - pure cyl | 0.33638 | 0.43651 | 0.5066 | 0.5867 | 0.73689 | 0.54 | 0.54
|  | Air flow for dilution mL/min | 1007.6654 | 1007.6654 | 1007.6654 | 1007.6654 | 1007.6654 | 1157 | 1157 | ng/min | 233
|  | Total flow mL/min | 1008.0017 | 1008.1019 | 1008.172 | 1008.2521 | 1008.4023 | 1157.54 | 1157.54 | ml/min | 1000
|  | C pure cylinder (ppbv) | C out (ppb) | C out (ppb) | C out (ppb) | C out (ppb) | C out (ppb) | C out (ppb) | C out (ppb) | M COS | 60.07
Bouteille n°4 | COS | 203700 | 67.977 | 88.2023 | 102.3576 | 118.5327 | 148.854 | 95.0274 | 118825 | C µg/m3 | 233
| DMS | 197500 | 65.908 | 85.5177 | 99.2421 | 114.9249 | 144.3233 | 92.135 | 115208.3333 | C ppb | 93.2465
| cycles repetitions (number) |  | 5 | 5 | 8 | 5 | 5
| duration of one cycle (s) |  | 600 | 600 | 600 | 600 | 600
| duration for one dilution point (s) |  | 3000 | 3000 | 4800 | 3000 | 3000
| total duration (s) | 16800
| total duration (hours) | 4:40 | 0:50 | 0:50 | 1:20 | 0:50 | 0:50
| Bouteille Chromatotec | [8 empty] | 95.0274
| Prévisionnel ChromaS: | gamme 0 à 400ppb | Point 1: 60
|  |  | Point 1 | Point 2 | Point 3 | Point 4 | Point 5 | Point 6
Set Point MFC N2
Set Point MFC Air to dilute
N2 flow mL/min - pure cyl
Air flow for dilution mL/min
|  | Total flow mL/min | 0 | 0 | 0 | 0 | 0 | 0
|  | C pure cylinder (ppv) | C out (ppb) | C out (ppb) | C out (ppb) | C out (ppb) | C out (ppb) | C out (ppb)
Bouteille n°1 | DES | 476 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| SO2 | 484 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| MeSH | 524 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| EtSH | 478 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| DMS | 484 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| CS2 | 488 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Bouteille n°2 | DMDS | 502 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Bouteille n°3 | H2S |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Bouteille n°4 | COS |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Bouteille H2S INERIS
| Prévisionnel ChromaS: | gamme 0 à 400ppb
|  |  | Point 1 | Point 2 | Point 3 | Point 4 | Point 5 | Point 6 | TEST
|  | Set Point MFC N2 | 36 | 70 | 50 | 28 | 20
|  | Set Point MFC Air to dilute | 26 | 60 | 65 | 65 | 80
| MFC 0-10mL/min | N2 flow mL/min - pure cyl | 3.6406 | 7.0449 | 5.0424 | 2.8396 | 2.0385 | 0.036 | 0.036
| MFC 0-1000mL/min | Air flow for dilution mL/min | 262.532 | 604.8906 | 655.2374 | 655.2374 | 806.278 | 0.72838 | 0.72838
|  | Total flow mL/min | 266.1726 | 611.9355 | 660.2798 | 658.077 | 808.3165 | 0.76438 | 0.76438
|  | C pure cylinder (ppb) | C out (ppb) | C out (ppb) | C out (ppb) | C out (ppb) | C out (ppb) | C out (ppb) | C out (ppb)
| H2S | 10000 | 136.7752 | 115.125 | 76.3671 | 43.1494 | 25.2196 | 470.9714 | 97.4155
| MeSH | 16400 | 224.3114 | 188.8051 | 125.242 | 70.765 | 41.3602 | 772.3931 | 159.7615
| DMS | 32100 | 439.0485 | 369.5514 | 245.1383 | 138.5096 | 80.9551 | 1511.8182 | 312.7038
| cycles repetitions (number) |  | 4 | 4 | 8 | 4 | 4
| duration of one cycle (s) |  | 600 | 600 | 600 | 600 | 600
| duration for one dilution point (s) |  | 2400 | 2400 | 4800 | 2400 | 2400
| total duration (s) | 14400
| total duration (hours) | 4:00 | 0:40 | 0:40 | 1:20 | 0:40 | 0:40

## Results_ChS_H2S_Ineris
Sampling date | Volume | AIR | H2S | METHYL-SH | DMS
Point 1
2022-07-12 11:37 | 2 | 359823.2 | 441245.8 | 544347.7 | 1246925.5
2022-07-12 11:47 | 2 | 358473.6 | 440513 | 525215.6 | 1226395.9
2022-07-12 11:57 | 2 | 354849.8 | 428024.1 | 528183 | 1225440.3
2022-07-12 12:07 | 2 | 357996.3 | 437374.5 | 536740.6 | 1244799
2022-07-12 12:17 | 2 | 359593.6 | 433738.8 | 538966.6 | 1242631.4
2022-07-12 12:27 | 2 | 356221.3 | 422946.7 | 542354.8 | 1234346
moy | 2 | 357826.3 | 433973.8167 | 535968.05 | 1236756.35
Point 2
2022-07-12 12:37 | 2 | 363689.7 | 314690.9 | 282976.8 | 633880.4
2022-07-12 12:47 | 2 | 361747.1 | 319003.3 | 282598.8 | 645157.8
2022-07-12 12:57 | 2 | 361847.1 | 313356.3 | 285029 | 632330.6
2022-07-12 13:07 | 2 | 360563.1 | 307357.3 | 288894.6 | 623511.8
2022-07-12 13:17 | 2 | 358731 | 309983 | 287760.2 | 636004.5
2022-07-12 13:27 | 2 | 360416.3 | 329258.1 | 283040.1 | 630087
2022-07-12 13:37 | 2 | 360384.7 | 315022.7 | 287606.4 | 641386.5
moy | 2 | 361054.1429 | 315524.5143 | 285415.1286 | 634622.6571
Point 3
2022-07-12 13:47 | 2 | 361748.1 | 104787.7 | 85271.7 | 207228.3
2022-07-12 13:57 | 2 | 362669.5 | 101973.7 | 88892.6 | 208188.9
2022-07-12 14:07 | 2 | 363839.1 | 105679.3 | 87677 | 212515.9
2022-07-12 14:17 | 2 | 362174.3 | 104005.2 | 90836.5 | 214235.7
2022-07-12 14:27 | 2 | 362459.5 | 95569 | 84696.4 | 205214.8
2022-07-12 14:37 | 2 | 360267.1 | 96149 | 82761.2 | 206385.8
2022-07-12 14:47 | 2 | 360687.1 | 92534.9 | 86211.9 | 200477.2
moy | 2 | 362016.1 | 99318.5167 | 86845.9333 | 207836.3833
Point 4
2022-07-12 15:07 | 2 | 366712.9 | 34969.5 | 29286.9 | 76079
2022-07-12 15:17 | 2 | 366972.2 | 37203 | 25112 | 75959.3
2022-07-12 15:27 | 2 | 368465.8 | 37590.3 | 25472 | 70173.8
2022-07-12 15:37 | 2 | 366288.9 | 34385.3 | 31828.6 | 71134.2
2022-07-12 15:47 | 2 | 367078.7 | 34534.4 | 30521.2 | 76437.8
moy | 2 | 367103.7 | 35736.5 | 28444.14 | 73956.82
Point 5
2022-07-12 16:27 | 2 | 367424.3 | 20398.3 | 17538.6 | 49201.4
2022-07-12 16:37 | 2 | 367833.3 | 20086.1 | 16488.9 | 43908.2
2022-07-12 16:47 | 2 | 369375.6 | 24080.2 | 18293.4 | 49815.1
2022-07-12 16:59 | 2 | 370306.2 | 22513.7 | 21603.5 | 44343.4
2022-07-12 17:09 | 2 | 369037.3 | 20275.6 | 21074.5 | 51698.6
2022-07-12 17:19 | 2 | 369457.6 | 20262.6 | 18720.6 | 49380.6
moy | 2 | 368905.7167 | 21269.4167 | 18953.25 | 48057.8833
Note on F86: Auteur: 02h16 somme de trois aires, manuellement car pic DMS large

## Results_ChS_COS_Ineris
Sampling date | Volume | AIR | COS | CS2
Point 1
2022-07-26 11:02 | 2 | 373152.6 | 172549.7 | 216863.8
2022-07-26 11:12 | 2 | 375238.3 | 167212.3 | 215648.7
2022-07-26 11:22 | 2 | 371028.2 | 179149 | 219227
moy |  | 373139.7 | 172970.3333 | 217246.5
Point 2
2022-07-26 11:42 | 2 | 373828 | 271968 | 339158
2022-07-26 11:52 | 2 | 373770.8 | 273686.5 | 338283
2022-07-26 12:02 | 2 | 376876.4 | 275527 | 340085.8
2022-07-26 12:12 | 2 | 372489.7 | 277423.3 | 334433
2022-07-26 12:22 | 2 | 376225.2 | 273569.2 | 339094.4
moy |  | 374638.02 | 274434.8 | 338210.84
Point 3
2022-07-26 13:02 | 2 | 376909.2 | 351122.9 | 447338.4
2022-07-26 13:12 | 2 | 373591.8 | 351810.4 | 460857.8
2022-07-26 13:22 | 2 | 375211 | 347348 | 445921.8
moy |  | 375237.3333 | 350093.7667 | 451372.6667
Point 4
2022-07-26 13:52 | 2 | 375191.7 | 408597.6 | 576930.1
2022-07-26 14:02 | 2 | 374815.5 | 405691.5 | 574165.5
2022-07-26 14:12 | 2 | 373828 | 406170.6 | 577927.4
2022-07-26 14:22 | 2 | 377327.9 | 404988.4 | 575902.9
moy |  | 375290.775 | 406362.025 | 576231.475
Point 5
2022-07-26 14:42 | 2 | 378868 | 476529.8 | 894426.1
2022-07-26 14:52 | 2 | 375140.3 | 479047 | 894234.9
2022-07-26 15:02 | 2 | 375045.8 | 473459 | 904890.3
2022-07-26 15:12 | 2 | 374577.8 | 474837.6 | 905397.9
2022-07-26 15:22 | 2 | 374795.5 | 475196.5 | 889296.2
moy |  | 375685.48 | 475813.98 | 897649.08

## ChS_Response_curves
Calibration curves adjusments - ChromaS
: 01 - BS set at 100 000 in ampli 1 ; BS = constant value
: 02 - cells to fill with the multiple-points calibration results
03 : Reponse factors obtained after calibration : C = a*(Area/BS ) ^b
CV = a*(Area/BS ) ^b
[8 empty] | a = | #VALUE!
DES | [7 empty] | b = | #VALUE!
Point n° | Vm (mol/L) | EXPECTED (ppb) | EXPECTED (mg/m3) | AREA | BS | MW (g) | AREA/BS
1 | 24.04 |  | 0 | #REF! | 100000 | 90.19 | #REF!
2 | 24.04 |  | 0 | #REF! | 100000 | 90.19 | #REF!
3 | 24.04 |  | 0 | #REF! | 100000 | 90.19 | #REF!
4 | 24.04 |  | 0 | #REF! | 100000 | 90.19 | #REF!
5 | 24.04 |  | 0 | #REF! | 100000 | 90.19 | #REF!
| 24.04 |  | 0 |  | 100000 | 90.19 | 0
| 24.04 |  | 0 |  | 100000 | 90.19 | 0
[8 empty] | a = | #VALUE!
DMDS | [7 empty] | b = | #VALUE!
Point n° | Vm (mol/L) | EXPECTED (ppb) | EXPECTED (mg/m3) | AREA | BS | MW (g) | AREA/BS
1 | 24.04 | - | #VALUE! | #REF! | 100000 | 94.2 | #REF!
2 | 24.04 | - | #VALUE! | #REF! | 100000 | 94.2 | #REF!
3 | 24.04 | - | #VALUE! | #REF! | 100000 | 94.2 | #REF!
4 | 24.04 | - | #VALUE! | #REF! | 100000 | 94.2 | #REF!
6 | 24.04 | - | #VALUE! | #REF! | 100000 | 94.2 | #REF!
| 24.04 |  | 0 |  | 100000 | 94.2 | 0
| 24.04 |  | 0 |  | 100000 | 94.2 | 0
[8 empty] | a = | 0.11595
H2S | [7 empty] | b = | 0.65022
Point n° | Vm (mol/L) | EXPECTED (ppb) | EXPECTED (mg/m3) | AREA | BS | MW (g) | AREA/BS
1 | 24.04 | 190.38 | 0.26989 | 433973.8167 | 100000 | 34.08 | 4.3397
2 | 24.04 | 152.66 | 0.21642 | 315524.5143 | 100000 | 34.08 | 3.1552
3 | 24.04 | 96.01 | 0.13611 | 99318.5167 | 100000 | 34.08 | 0.99319
4 | 24.04 | 48.69 | 0.069025 | 35736.5 | 100000 | 34.08 | 0.35736
5 | 24.04 | 24.67 | 0.034973 | 21269.4167 | 100000 | 34.08 | 0.21269
6 | 24.04 |  | 0 |  | 100000 | 34.08 | 0
| 24.04 |  | 0 |  | 100000 | 34.08 | 0
[8 empty] | a = | 0.10628
COS | [7 empty] | b = | 0.73361
Point n° | Vm (mol/L) | EXPECTED (ppb) | EXPECTED (mg/m3) | AREA | BS | MW (g) | AREA/BS
1 | 24.04 | 66.38 | 0.16587 | 172970.3333 | 100000 | 60.07 | 1.7297
2 | 24.04 | 86.12 | 0.21519 | 274434.8 | 100000 | 60.07 | 2.7443
3 | 24.04 | 99.95 | 0.24975 | 350093.7667 | 100000 | 60.07 | 3.5009
4 | 24.04 | 115.74 | 0.28921 | 406362.025 | 100000 | 60.07 | 4.0636
5 | 24.04 | 145.35 | 0.36319 | 475813.98 | 100000 | 60.07 | 4.7581
6 | 24.04 |  | 0 |  | 100000 | 60.07 | 0
| 24.04 |  | 0 |  | 100000 | 60.07 | 0
[8 empty] | a = | #VALUE!
SO2 | [7 empty] | b = | #VALUE!
Point n° | Vm (mol/L) | EXPECTED (ppb) | EXPECTED (mg/m3) | AREA | BS | MW (g) | AREA/BS
1 | 24.04 |  | 0 | #REF! | 100000 | 64.07 | #REF!
2 | 24.04 |  | 0 | #REF! | 100000 | 64.07 | #REF!
3 | 24.04 |  | 0 | #REF! | 100000 | 64.07 | #REF!
4 | 24.04 |  | 0 | #REF! | 100000 | 64.07 | #REF!
5 | 24.04 |  | 0 | #REF! | 100000 | 64.07 | #REF!
6 | 24.04 |  | 0 |  | 100000 | 64.07 | 0
| 24.04 |  | 0 |  | 100000 | 64.07 | 0
[8 empty] | a = | 0.28373
MeSH | [7 empty] | b = | 0.63046
Point n° | Vm (mol/L) | EXPECTED (ppb) | EXPECTED (mg/m3) | AREA | BS | MW (g) | AREA/BS
1 | 24.04 | 303.42 | 0.60722 | Sature | 100000 | 48.11 | #VALUE!
2 | 24.04 | 243.31 | 0.48692 | 285415.1286 | 100000 | 48.11 | 2.8542
3 | 24.04 | 153.01 | 0.30621 | 86845.9333 | 100000 | 48.11 | 0.86846
4 | 24.04 | 77.6 | 0.1553 | 28444.14 | 100000 | 48.11 | 0.28444
5 | 24.04 | 39.32 | 0.078689 | 18953.25 | 100000 | 48.11 | 0.18953
6 | 24.04 |  | 0 |  | 100000 | 48.11 | 0
| 24.04 |  | 0 |  | 100000 | 48.11 | 0
[8 empty] | a = | #VALUE!
EtSH | [7 empty] | b = | #VALUE!
Point n° | Vm (mol/L) | EXPECTED (ppb) | EXPECTED (mg/m3) | AREA | BS | MW (g) | AREA/BS
1 | 24.04 |  | 0 | #REF! | 100000 | 62.13 | #REF!
2 | 24.04 |  | 0 | #REF! | 100000 | 62.13 | #REF!
3 | 24.04 |  | 0 | #REF! | 100000 | 62.13 | #REF!
4 | 24.04 |  | 0 | #REF! | 100000 | 62.13 | #REF!
5 | 24.04 |  | 0 | #REF! | 100000 | 62.13 | #REF!
6 | 24.04 |  | 0 |  | 100000 | 62.13 | 0
| 24.04 |  | 0 |  | 100000 | 62.13 | 0
[8 empty] | a = | 0.26457
DMS | [7 empty] | b = | 0.66958
Point n° | Vm (mol/L) | EXPECTED (ppb) | EXPECTED (mg/m3) | AREA | BS | MW (g) | AREA/BS
1 | 24.04 | 389.69 | 1.0071 | Sature | 100000 | 62.13 | #VALUE!
2 | 24.04 | 312.48 | 0.80759 | 634622.6571 | 100000 | 62.13 | 6.3462
3 | 24.04 | 196.51 | 0.50787 | 207836.3833 | 100000 | 62.13 | 2.0784
4 | 24.04 | 99.67 | 0.25759 | 73956.82 | 100000 | 62.13 | 0.73957
5 | 24.04 | 50.49 | 0.13049 | 48057.8833 | 100000 | 62.13 | 0.48058
6 | 24.04 |  | 0 |  | 100000 | 62.13 | 0
| 24.04 |  | 0 |  | 100000 | 62.13 | 0
[8 empty] | a = | 0.13885
CS2 | [7 empty] | b = | 0.55269
Point n° | Vm (mol/L) | EXPECTED (ppb) | EXPECTED (mg/m3) | AREA | BS | MW (g) | AREA/BS
1 | 24.04 | 66.81 | 0.2116 | 217246.5 | 100000 | 76.14 | 2.1725
2 | 24.04 | 86.69 | 0.27457 | 338210.84 | 100000 | 76.14 | 3.3821
3 | 24.04 | 100.6 | 0.31862 | 451372.6667 | 100000 | 76.14 | 4.5137
4 | 24.04 | 116.5 | 0.36898 | 576231.475 | 100000 | 76.14 | 5.7623
5 | 24.04 | 146.3 | 0.46336 | 897649.08 | 100000 | 76.14 | 8.9765
6 | 24.04 |  | 0 |  | 100000 | 76.14 | 0
| 24.04 |  | 0 |  | 100000 | 76.14 | 0
Note on F37: Auteur: Légère saturation
Note on C105: Auteur: Peak découpé - pas pris en compte

## ChS_reponse_formulas
Reponse factors obtained after calibration. To copy/paste in the Vistachrom substance tables
C = a*(Area/BS ) ^b
LDl calculation
| Ampli 2 |  | Ampli 1 | A3 | Ampli 3 |  | M(g/mol) | BS | Area mini | LDL Ampli 2 (mg/m3) |  | LDL Ampli 2 (ppb)
| a | b | a | b | a | b |  |  | 1000
DES | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 90.19 | 100000 | 1000 | #VALUE! | DES | #VALUE!
DMDS | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 94.2 | 100000 | 1000 | #VALUE! | DMDS | #VALUE!
H2S | 0.11595 | 0.65022 | 0.51818 | 0.65022 | 0.025944 | 0.65022 | 34.08 | 100000 | 1000 | 0.0058052 | H2S | 4.095
COS | 0.10628 | 0.73361 | 0.57554 | 0.73361 | 0.019627 | 0.73361 | 60.07 | 100000 | 1000 | 0.0036245 | COS | 1.4505
SO2 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 64.07 | 100000 | 1000 | #VALUE! | SO2 | #VALUE!
MeSH | 0.28373 | 0.63046 | 1.2116 | 0.63046 | 0.066444 | 0.63046 | 48.11 | 100000 | 1000 | 0.01556 | MeSH | 7.775
EtSH | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | 62.13 | 100000 | 1000 | #VALUE! | EtSH | #VALUE!
DMS | 0.26457 | 0.66958 | 1.2363 | 0.66958 | 0.056619 | 0.66958 | 62.13 | 100000 | 1000 | 0.012117 | DMS | 4.6883
CS2 | 0.13885 | 0.55269 | 0.49573 | 0.55269 | 0.038893 | 0.55269 | 76.14 | 100000 | 1000 | 0.010894 | CS2 | 3.4395
CS2 | 0.13885 | 0.55269 | 0.49573 | 0.55269 | 0.038893 | 0.55269 | 76.139 | 100000 | 1000 | 0.010894 | CS2 | 3.4396